## Slide 1
Winter in Mazandaran forest, IRAN
Max E(Nuf) QC Model
Incorporating frequency into QC planning
Hassan Bayat
DCLS
Sina Clinical Laboratory, Qaem Shahr (Iran)

## Slide 2
QC frequency or Run size?
ISO 15189:
QC frequency … “based on…the risk of harm to the patient from an erroneous result”
Background
Conventional QC Planning:
Set Performance Specification; TEa
Determine Performance Characteristics; Bias & CV
SQC Characteristics; Ped & Pfr

## Slide 3
Max E(Nuf) SQC Model
Developed by Dr. C. Parvin
Provides scientific basis for QC frequency
Incorporates QC frequency into QC planning
Endorsed by CLSI C24-Ed4 (2016)
QC frequency or Run size?
ISO 15189:
QC frequency … “based on…the risk of harm to the patient from an erroneous result”
Background
Conventional QC Planning:
Set Performance Specification; TEa
Determine Performance Characteristics; Bias & CV
SQC Characteristics; Ped & Pfr
Mitigating patient risk to tolerable limits

## Slide 4
Chart: 
 | -0.0062335 | 0.14337
 | -0.0062335 | 0.1309
 | -0.0062335 | 0.1309
 | 0.074802 | 0.074802
Mean | -0.0062335 | 0.1309
Max E(Nuf) model is based on Total Analytical Error model
Incorrect results
Correct results
Incorrect results
SD
SD

## Slide 5
Chart: 
 | -0.0062335 | 0.14337
 | -0.0062335 | 0.1309
 | -0.0062335 | 0.1309
 | 0.074802 | 0.074802
Mean | -0.0062335 | 0.1309
Incorrect results
Correct results
Incorrect results
Bias
Max E(Nuf) model is based on Total Analytical Error model

## Slide 6
Chart: 
 | -0.0046389 | 0.10669
 | -0.0046389 | 0.097416
 | -0.0046389 | 0.097416
 | 0.055666 | 0.055666
Mean | -0.0046389 | 0.097416
Incorrect results
Correct results
Incorrect results
Related to Sigma Metric
SM = (TEa – B) / SD
Bias
Max E(Nuf) model is based on Total Analytical Error model

## Slide 7
Target
-TEa
+TEa
QC
QC
QC
QC
QC
QC
Schematic view of Max E(Nuf) model
Error Occurrence to Error Detection
How many increased errors are produced?
Of the them, how many are reported?
Cor.
SHIFT

## Slide 8: E(Nuf): Expected Number of Increased Unreliable Final Results
Unreliable: Incorrect or Erroneous
Bracketed QC
Final
QC
QC
U
QC
I
I
U
?
QC
QC
QC
I
U
I
U
I
I
QC
∆PE = Stable PE – Unstable PE
U
Reported or Finalized
Not reported or
Not Finalized
U
QC
Unstable PE
Stable PE
PE: Probability of Error production
E(Nuf)
For a given size of error and QC: Increased errors reported from shift to detection

## Slide 9
∆PE Increase in the probability of producing error
      Shift size; The larger is shift, the higher is ∆PE
Factors contributing to E(Nuf)

## Slide 10
∆PE Increase in the probability of producing error
      Shift size; The larger is shift, the higher is ∆PE
      Stable Bias
Factors contributing to E(Nuf)

## Slide 11
∆PE Increase in the probability of producing error
      Shift size; The larger is shift, the higher is ∆PE
      Stable Bias
      Stable Imprecision
Factors contributing to E(Nuf)
      The higher is Sigma, the lower ∆PE
SM = (TEa – B) / SD

## Slide 12: Factors contributing to E(Nuf)
∆PE
ARLed = 1/Ped
ARLed: Average Run Length
Ped: Probability of Error Detection
Chart: Power Graph
1:4s, N1
1:2s, N2
ARLed
1/0.90 = 1.1
0.08
ARLed
1/0.08 = 12.5
0.90

## Slide 13: Factors contributing to E(Nuf)
∆PE
ARLed
Chart: Power Graph
Chart: SE
PAccept = 1 - Ped
ARLed: Average Run Length
Ped: Probability of Error Detection
Ped = 15%
Pacpt
1 – 0.15
= 85%

## Slide 14: Factors contributing to E(Nuf)
∆PE
ARLed
Run Length; M Number of patient samples between QC events

## Slide 15: Factors contributing to E(Nuf)
∆PE
ARLed
Run Length; M
QC
QC
QC
QC
QC
QC
QC
`
QC

## Slide 16
∆PE
ARLed
Run Length; M
Shift point
QC
QC
Factors contributing to E(Nuf)

## Slide 17
∆PE
ARLed
Run Length; M
Shift point
QC
QC
M/2
Factors contributing to E(Nuf)

## Slide 18
ERROR SIZE
WASTING RATE
WASTE
What about “Max”

## Slide 19
ERROR SIZE
Probability of missing the problem
Neglecting

## Slide 20
Max Waste
ERROR SIZE
Probability
Wasting
Neglecting

## Slide 21
ERROR SIZE
Max Waste
Worst Scenario
Wasted Water Volume

## Slide 22
Chart: SE
Example: Sigma = 5 QC: 1:3s, N1, M200

## Slide 23
Chart: SE
Error production
ΔPE
Example: Sigma = 5; QC: 1:3s N1 M200

## Slide 24
Chart: SE
1:3s, N1
PAcceptance
1 - Ped
5σ Error production
ΔPE
Example: Sigma = 5; QC: 1:3s N1 M200

## Slide 25
Example
Sigma = 5 QC: 1:3s, N1, M200
SE (x SD) | ΔPE | ARLed | E(Nuf)
0.5 | 0.0003% | 155 | 0.1
1 | 0.003% | 44 | 0.3
3 | 2% | 2 | 3.4
3.5 | 6.7% | 1.4 | 3.9
4 | 16% | 1.2 | 3.5
7 | 98% | 1.003 | 0.003
8 | 99.8% | 1.0000004 | 0.00003
PE-Stable = 0.00006%
Max E(Nuf)
Chart: SE
Paccept
for 1:3s N1
∆PE
for SM=5
Worst Scenario

## Slide 26
Chart: SE (multiples of SD)
Chart: SE
Paccept
for 1:3s N1
∆PE
for SM=5
Max E(Nuf) = 3.9
3.5
SE (x SD) | ΔPE | ARLed | MaxE (Nuf)
3.5 | 6.77% | 1.4 | 3.9
Patient Risk Factor; PRF
Patient Risk Factor
Example:
Sigma = 5 QC Procedure: 1:3s N1 M200

## Slide 27: Allowable PRF or allowable Max E(Nuf)
Considerations:
Probability of hazard to patient
  Higher hazard → Lower PRF
Interpreted standalone or combined with other tests
  Standalone→ Lower PRF
Speed of acting upon an erroneous result
  Faster use → Lower PRF
The stability of the system
  Less stable → Lower PRF
Sigma Metric
  Higher Sigma → Higher PRF
No consensus at present

## Slide 28
Yago M, Alcover S.
Clin Chem62:7; 2016
Nomograms for M = 100

## Slide 29
CCLM; 55(11), 23017
MR3 N3: 13s/2of32s/R4s/31s
MR3 N6: 13s/2of32s/R4s/31s/6x
MR2 N2: 13s/22s/R4s
MR4 N4: 13s/22s/R4s/41s
Max E(Nuf)/Run Size Ratio

## Slide 30
Planning Risk-Based Statistical Quality Control Strategies: Graphical Tools to Support the New Clinical and Laboratory Standards Institute C24-Ed4 Guidance
Bayat H, Westgard SA, Westgard JO
JALM; September 2017
Selecting a Risk-Based SQC Procedure for a HbA1c Total QC Plan
Westgard SA, Bayat H, Westgard JO
Journal of Diabetes Science and Technology 2017
Planning Risk-based SQC Schedules for Bracketed Operation of Continuous Production Analyzers
Westgard JO, Bayat H, Westgard SA
Clinical Chemistry Feb 2018, 64 (2) 289-296

## Slide 31
Candidate SQC (Pfr) | Maximum Run Size / Ped |  |  |  |  | 
 | 6-Sigma | 5.5-Sigma | 5.0-Sigma | 4.5-Sigma | 4.0-Sigma | 3.5-Sigma
MR N4 (0.03) | ≥1000 | ≥1000 | ≥1000 | 800 | 190 | 40
 | 1.00 | 1.00 | 1.00 | 0.98 | 0.91 | 0.73
1:3s N4 (0.01) | ≥1000 | ≥1000 | 900 | 220 | 60 | 15
 | 1.00 | 1.00 | 0.99 | 0.89 | 0.66 | 0.41
MR N2 (0.01) | ≥1000 | ≥1000 | 470 | 120 | 40 | 10
 | 1.00 | 0.99 | 0.94 | 0.82 | 0.59 | 0.35
1:2s N1 (0.05) | ≥1000 | ≥1000 | 370 | 140 | 50 | 20
 | 0.98 | 0.97 | 0.91 | 0.80 | 0.62 | 0.43
1:3s N2 (0.00) | ≥1000 | 700 | 220 | 70 | 25 | <10
 | 0.98 | 0.96 | 0.86 | 0.70 | 0.48 | 0.25
1:2.5s N1 (0.01) | 1000 | 400 | 150 | 50 | 20 | <10
 | 0.94 | 0.92 | 0.82 | 0.66 | 0.44 | 0.24
1:3s N1 (0.00) | 380 | 150 | 50 | 20 | 10 | <10
 | 0.85 | 0.82 | 0.67 | 0.46 | 0.27 | 0.11
Sigma Run Size Matrix
Establishing Evidence-Based Statistical Quality Control Practice
Westgard JO, Westgard SA
Am J Cling Pathol 2018
Data
SQC
13s
22s
R4s
41s
6X
Take Corrective Action
Report Results
No
6σ
5σ
4σ
3σ
No
No
No
Yes
Yes
Yes
Yes
Yes
N=2
Run
Size
1000
N=2
Run
Size
450
N=4
Run
Size
200
N=6
Run
Size
45
No
Westgard Sigma Rules® with Run Sizes

## Slide 32
https://www.westgard.com/qc-frequency-calculator.htm

## Slide 33
https://tools.westgard.com/frequency_calculator.shtml

## Slide 34
https://tools.westgard.com/frequency_calculator.shtml

## Slide 36
https://tools.westgard.com/frequency_calculator.shtml

## Slide 37
عیار سیگمای بالای 6، در محاسبه برابر 6 در نظر گرفته میشود

## Slide 38
عیار سیگمای بالای 6، در محاسبه برابر 6 در نظر گرفته میشود

## Slide 39
عیار سیگمای بالای 6، در محاسبه برابر 6 در نظر گرفته میشود

## Slide 40
عیار سیگمای بالای 6، در محاسبه برابر 6 در نظر گرفته میشود

## Slide 41
عیار سیگمای بالای 6، در محاسبه برابر 6 در نظر گرفته میشود
D
E
C
R
E
A
S
E

## Slide 42
عیار سیگمای بالای 6، در محاسبه برابر 6 در نظر گرفته میشود
تناوب انجام پ ک
Frequency
اندازه دود
Run size

## Slide 43
EXAMPLE

## Slide 44
D
E
C
R
E
A
S
E

## Slide 50
Hb A1C
%Hb
%Hb
Hb A1C
EXAMPLE

## Slide 51
Sigma: 5.45
Patient Risk: 1
Sigma: 4.29
Patient Risk: 1

## Slide 52
%Hb
%Hb
Example:
4
Bio. Test X
Bio. Test Y

## Slide 53
Sigma: 5.45
Patient Risk: 1
Sigma: 4.29
Patient Risk: 4
1:3.5s N=2
M=200
1:3s N=2
M=200
1:3s N=1
M=135
1:2.5s N=1
M=135
Patient Risk: 1

## Slide 54
https://tools.westgard.com/frequency_calculator2.shtml

## Slide 55
https://tools.westgard.com/frequency_calculator3.shtml.

## Slide 57
Restricted to 3s:
Run size goes more than 1000
Tacking large time to complete a run
Many repeats following corrective actions
Impractical large run sizes for High Sigma performances

## Slide 59
THANKS!