## Slide 1: CLSI EP15-A3 (Ed4)
Hassan Bayat
Clinical Laboratory Scientist
Sina Clinical Laboratory (Iran)

## Slide 4: Validation vs Verification
Validation: Evaluating method
Verification: Evaluating user lab’s performance
Long-term studies
Meeting intended purposes
Short-term studies
Consistency with manufacturers claims
Precision validation: CLSI EP05
Bias validation: CLSI EP09
Precision validation: CLSI EP15

## Slide 6: Application of EP15
Main application: Implementing an established method
Other: Confirming performance following corrective actions

## Slide 7: Overview
5 days (not necessarily consecutive)
If test materials with known target values, estimate bias in addition to verifying imprecision
Training operators
Material selection
Implementing QC
Verifying IMPRECISION
Estimating BIAS
END

## Slide 8: Run definition in EP15
The time in which stable performance is expected
Not more than 24 hours
Not more than the duration specified by manufacture
Consider manufacture’s validation study

## Slide 9: Specify APS (analytical performance specification)
Set APS Before choosing a method
If manufacture’s claim doesn’t meet APS:
Choose another method
Or
More lenient APS

## Slide 10: Documentation
The study as a whole as well the details of runs and replicates
Instrument, measurand (analyte, quantity, unit)
Reagent/calibrator lot number
Rationale for selecting the levels
Test material structure and, if available, its assigned value
Software
Operators, data evaluation and interpretation
Study design; decision for extending study

## Slide 11
Part 1

## Slide 12
Repeatability
Within-Lab
Reproducibility
Run-to-run
Day-to-day

## Slide 13
Repeatability
Within-Lab
Reproducibility
Run-to-run
Day-to-day
Within-run
precision
One lab
Short time
One operator
Same instrument
Same devices

## Slide 14
Within-run
precision
One lab
Short time
One operator
Same instrument
Same devices
One lab
Same instrument
Specified time
Different operators
Different devices
May different reagent/calibrator lots
Total precision
Repeatability
Within-Lab
Reproducibility
Run-to-run
Day-to-day

## Slide 15
Different labs
Long time
Different reagent lots
Different calibrator lots
Different operators
Different instruments/devices
Repeatability
Within-Lab
Reproducibility
Run-to-run
Day-to-day
One lab
Same instrument
Specified time
Different operators
Different devices
May different reagent/calibrator lots
Total precısıon
Within-run
precision
One lab
Short time
One operator
Same instrument
Same devices

## Slide 16
Different labs
Long time
Different reagent lots
Different calibrator lots
Different operators
Different instruments/devices
Repeatability
Within-Lab
Reproducibility
Run-to-run
Day-to-day
One lab
Short time
One operator
Same instrument
Same devices
One lab
Same instrument
Specified time
Different operators
Different devices
May different reagent/calibrator lots

## Slide 18: Overview, Limitations
At least 5 days, 1 run per day, 5 replicates per run
False rejection = 5%
Detecting large deviations from manufacture’s claim
Claims are treated as if they have no uncertainty

## Slide 19
Compromise between lab resources and study’s quality
Evaluating repeatability & within-lab precision at the same time
Robustness of calculations and estimates
Material
Time
Sufficient evaluation
Informative evaluation
Advantages

## Slide 20
Precısıon type | Manufacture claim | Lab estimate
Repeatability | σR | SR
Within-lab | σWL | SWL
Population
Sample
Symbols

## Slide 21: Samples
Type:
Preferably patient specimen (single or pooled)
Commercial specimens
Be compatible with manufacture study

## Slide 22
Level
Clinical decision levels
Normal & abnormal intervals
Levels close to manufacture study makes easier to verify claims
Not out of manufacture claimed levels

## Slide 23: Assaying samples
Calibrate at the start at the study
If manufacture has done several calibrations, consider extra calibrations
All the users be involved; even for fully automated analyzers

## Slide 24: Data integrity
Chart: 
| 1 | 2 | 3 | 4 | 5
 | 140 | 139 | 138 | 138 | 140
 | 140 | 143 | 141 | 143 | 137
 | 140 | 138 | 136 | 141 | 136
 | 141 | 144 | 142 | 143 | 144
 | 139 | 140 | 141 | 138 | 141


 | 140.12 | 140.12
 | 139 | 139
 | 140.8 | 140.8
 | 138.2 | 138.2
 | 142.8 | 142.8
Identifying results or runs that are different
Evaluate results each day → Aberrant results not reflecting performance

## Slide 25: Data integrity
Chart: 
| 1 | 2 | 3 | 4 | 5
 | 140 | 139 | 138 | 138 | 140
 | 140 | 143 | 141 | 143 | 137
 | 140 | 138 | 136 | 141 | 136
 | 141 | 144 | 142 | 143 | 144
 | 139 | 140 | 141 | 138 | 141


 | 140.12 | 140.12
 | 139 | 139
 | 140.8 | 140.8
 | 138.2 | 138.2
 | 142.8 | 142.8
Identifying results or runs that are different
Evaluate results each day → Aberrant results not reflecting performance

## Slide 26: Data integrity
Evaluate results each day → Aberrant results not reflecting performance
Don’t rely on guess and assumption
Don’t delete or repeat results just for being different
Correcting typing mistakes, misplaced samples, …
Irrelevant errors: clog, insufficient volume, …
If QC failed, delete whole run
Document any correction, deletion, missing, and repeats

## Slide 27
DON’T delete any result as statistical “outlier” at this phase
After correcting/repeating attributable results; if still aberrant results:
Do the calculations with all results; If claim is verified, outlier test isn’t needed
If not:
Do statistical outlier test and repeat calculations after deleting probable statistical outlier

## Slide 28
Delete or keep?
Maybe because of an unknown irrelevant reason to performance
Maybe because of performance; inconsistent with the small number of data but consistent with a larger study
Treating Probable outliers (after deleting attributable results )

## Slide 29
Chart: 
| 47 | 48 | 49 | 50 | 51 | 52
 | 220 | 221 | 220 | 221 | 226 | 221
Chart: 
Example: Inconsistent with a small number of data but consistent with a larger study

## Slide 30
Chart: 
| 47 | 48 | 49 | 50 | 51 | 52
 | 220 | 221 | 220 | 221 | 226 | 221
Chart: 
Example: Inconsistent with a small number of data but consistent with a larger study

## Slide 31
Chart: 
| 47 | 48 | 49 | 50 | 51 | 52
 | 220 | 221 | 220 | 221 | 226 | 221
Chart: 
Example: Inconsistent with a small number of data but consistent with a larger study

## Slide 32
Chart: 
| 47 | 48 | 49 | 50 | 51 | 52
 | 220 | 221 | 220 | 221 | 226 | 221
Chart: 
Keeping: SD = 2.06
Deleting: SD = 0.49
Better estimate
Example: Inconsistent with a small number of data but consistent with a larger study

## Slide 33
Risk of false rejection
Risk of optimistic estimation
Keeping while irrelevant to performance
Delete while truly related to performance
Confirm an apparently outlier with an appropriate statistical test; e.g. Grubbs’ test.
Delete at most 1 result per level, and 2 results per whole study.
Note: Even if a results is confirmed as a statistical outlier but doesn’t lead to fail verification, don’t delete it.

## Slide 34: Grubbs’ test
Calculate grand mean, , and SD from all results;
Determine G from table
G factor for 5 runs | 
G | N
3.087 | 23
3.112 | 24
3.135 | 25
A result is outlier if, and only if, it more the G•SD away from mean

## Slide 35
EXAMPLE
= 140.1
SD = 2.30
Grubs’ limits = ± G*SD
= 140.1 ± (3.135 x 2.30) = 132.9 & 147.3


G factor for 5 runs | 
G | N
3.087 | 23
3.112 | 24
3.135 | 25

## Slide 36
Number of runs adjusted for missing data
Effective number of runs, n0
Effective number of runs (5 runs) | 
n0 | N
4.565 | 23
4.792 | 24
5 | 25

## Slide 37: Estimating lab precision; ANOVA
Separating total variation to “within-group” and “between-group” components
within-run & between-run precision

## Slide 38: Estimating lab precision; ANOVA
nRun - 1
N - nRun
N - 1
Sum Square (of Deviations from Mean)
Degrees of Freedom
Mean Square
(SS/DF)
Separating total variation to “within-group” and “between-group” components
within-run & between-run precision

## Slide 39: Estimating lab precision; ANOVA
Separating total variation to “within-group” and “between-group” components
within-run & between-run
nRun - 1
N - nRun
N - 1
Sum Square (of Deviations from Mean)
Degrees of Freedom
Mean Square
(SS/DF)
VR = MS2
VB = (MS1 – MS2)/n0
Repeatability Variance
“Pure” between-run variance
VWL = VW + VB
Within-lab variance

## Slide 40
VR = MS2
VB = (MS1 – MS2)/n0
Repeatability Variance
“Pure” between-run variance
VWL = VW + VB
Within-lab variance
SR =
SB =
SWL =
%CVR = (SR /
%CVWL = (SWL /
%CVB = (SB /
Having components help specify sources of variation

## Slide 41
Chart: n
| 1 | 2 | 3 | 4 | 5
 | 140 | 139 | 138 | 138 | 140
 | 140 | 143 | 141 | 143 | 137
 | 140 | 138 | 136 | 141 | 136
 | 141 | 144 | 142 | 143 | 144
 | 139 | 140 | 141 | 138 | 141


 | 140.12 | 140.12
 | 139 | 139
 | 140.8 | 140.8
 | 138.2 | 138.2
 | 142.8 | 142.8
SUMMARY |  |  |  | 
Groups | Count | Sum | Average | Variance
Run 1 | 5 | 695 | 139 | 1
Run 2 | 5 | 704 | 140.8 | 6.2
Run 3 | 5 | 691 | 138.2 | 5.2
Run 4 | 5 | 714 | 142.8 | 1.7
Run 5 | 5 | 699 | 139.8 | 1.7
ANOVA |  |  | 
Source of Variation | SS | DF | MS
Between Run | 63.44 | 4 | 15.86
Within Run | 63.20 | 20 | 3.16
Total | 126.64 | 24 | 
VR = 3.16
SR = 1.78
VB = (15.86 – 3.16)/5 = 2.54
SB = 1.59
VWL = 3.16 + 2.54 = 5.70
SWL = 2.39
CVR = 1.27%
CVB = 1.13%
CVWL = 1.71%
(S/140.1) * 100

## Slide 42: Determine manufacture claim
If lab mean differs from manufacture’s mean (that usually is):
Claimed CV at closest PI mean to lab mean
Interpolation or averaging 2 or 3 claimed CVs around lab mean
Using precision profile (if present)

## Slide 43: Compare estimates to claims
If lab performance is the same as manufacture and repeats verification study many times:
Approximately half of estimates would be less than claim
Approximately half of estimates would be higher than claim

## Slide 44: Compare estimates to claims
Upper Verification Limit; UVL
(for each estimate)False rejection = 50%
σ
Pass
Fail
(for whole study)FR = 1 – (50%)2*nSam
If lab performance is the same as manufacture and repeats verification study many times:
Approximately half of estimates would be less than claim
Approximately half of estimates would be higher than claim
nSam = 1 → FR = 75%
nSam = 2 → FR = 94%
nSam = 1 → FR = 98%

## Slide 45
Upper Verification Limit; UVL
σ
UVL = F * σ
(for whole study) False rejection = 5%

## Slide 46: Calculate F (thereby UVL)
Rule of thumb: F = 1.3
To calculate F exactly:
Determine dfR and dfWL
→ UVL = 1.3 * σ
dfR = N – nRun
dfWL
ρ = σWL / σR

## Slide 47: Calculate F (thereby UVL)
Rule of thumb: F = 1.3
To calculate F exactly:
Determine dfR and dfWL
→ UVL = 1.3 * σ
dfR = N – nRun
dfWL
ρ = σWL / σR
Find F from table for given df and nSamp

## Slide 48: Summary of Calculations
SR
UVLR
ANOVA
σR
σWL
dfR
dfWL
nSam
FR
FWL
UVLWL
At least 5 day
5 replicate/day
SWL
F * σ
Package insert

## Slide 49: Summary of Calculations
σR
σWL
F * σ
ANOVA
dfR
dfWL
nSam
FR
FWL
At least 5 day
5 replicate/day
Package insert
Comparing estimates to claims:
Compare to claim; if not passed,
Compare to UVL
SR
UVLR
UVLWL
SWL
σR
σWL
NOTE: Because of small sample volume, both false verification and false rejection is probable

## Slide 50: Probable scenarios
1A) All estimates are smaller than claim or well smaller than UVL
The overall study is consistent with the claims
Imprecision in the user’s laboratory is comparable to the claims
1B) All passed, but some pass just barely
The overall study is consistent with claims
Monitor daily QC carefully to ensure imprecision remains consistent with claim
Interpretation

## Slide 51
2) Most estimates accepted but some fail
The lab director should decide considering clinical aspect
(if decided to use method) Imprecision exhibited by daily QC be carefully monitored
3) Several estimates fail
The study as a whole is not consistent with the claims
Reject the method
Troubleshoot and repeat study
Perform CLSI EP05

## Slide 52: EXAMPLE: Verifying a Ferritin assay
Manual assay → Fully automate assay; AMR = 10-1000 μg/L
Precision is concentration-depended
Lower, higher & normal results are clinically important
Manufacture study included several calibrations
Levels: Low, Normal, and High
Sample: Pooled sera
Calibration: Before each run

## Slide 53
Chart: Repeat
| 1 | 2 | 3 | 4 | 5
 | 140 | 139 | 138 | 138 | 140
 | 140 | 143 | 141 | 143 | 137
 | 140 | 138 | 136 | 141 | 136
 | 141 | 144 | 142 | 143 | 144
 | 139 | 140 | 141 | 138 | 141


 | 140.12 | 140.12
 | 139 | 139
 | 140.8 | 140.8
 | 138.2 | 138.2
 | 142.8 | 142.8
Chart: Repeat
| 1 | 2 | 3 | 4 | 5
 | 26.6 | 25.2 | 30.2 | 27.6 | 25.6
 | 24.3 | 25.7 | 23.8 | 25.3 | 24.1
 | 26.1 | 24 | 25.4 | 26 | 24.3
 | 26.5 | 27.1 | 25.9 | 25.5 | 25.5
 | 24.5 | 26.4 | 25.8 | 26 | 25.1


 | 25.7 | 25.7
 | 27.04 | 27.04
 | 24.64 | 24.64
 | 25.16 | 25.16
 | 26.1 | 26.1
Chart: Repeat
| 1 | 2 | 3 | 4 | 5
 | 606 | 627 | 621 | 606 | 620
 | 612 | 610 | 611 | 595 | 630
 | 649 | 626 | 636 | 639 | 648
 | 615 | 633 | 605 | 616 | 625
 | 622 | 632 | 646 | 619 | 623


 | 622.88 | 622.88
 | 616 | 616
 | 611.6 | 611.6
 | 639.6 | 639.6
 | 618.8 | 618.8

## Slide 54
G factor for 5 runs | 
G | N
3.087 | 23
3.112 | 24
3.135 | 25
Grubbs’ limits = (G • SD)
Grubbs’ Test

## Slide 55
Grubbs’ Test
G factor for 5 runs | 
G | N
3.087 | 23
3.112 | 24
3.135 | 25
Grubbs’ limits = (G • SD)

## Slide 56
ANOVA
VW (VR) = MS2
VB = (MS1 – MS2)/n0
VWL = VW + VB
Effective number of runs (5 runs) | 
n0 | N
4.565 | 23
4.792 | 24
5 | 25
VB = (4.238 – 1.3284)/5
VB = 0.58192
VWL = 1.3284 + 0.58192
VWL = 0.58192
VW = 1.3284

## Slide 57
ANOVA
VW (VR) = MS2
VB = (MS1 – MS2)/n0
VWL = VW + VB
Effective number of runs (5 runs) | 
n0 | N
4.565 | 23
4.792 | 24
5 | 25
VB = (4.238 – 1.3284)/5
VB = 0.58192
VWL = 1.3284 + 0.58192
VWL = 0.58192
VW = 1.3284

## Slide 58
Claims & UVLs

## Slide 59
Claims & UVLs
dfR = N – nRun
dfR = 25 – 5 = 20

## Slide 60
Claims & UVLs
dfR = N – nRun
dfR = 25 – 5 = 20

## Slide 61
Claims & UVLs
dfR = N – nRun
dfR = 25 – 5 = 20

## Slide 62
Claims & UVLs
ρ = σWL / σR

## Slide 63
Claims & UVLs
ρ = σWL / σR

## Slide 64
Claims & UVLs
ρ = σWL / σR

## Slide 65
Claims & UVLs
ρ = σWL / σR

## Slide 66
Claims & UVLs

## Slide 67
Comparing Estimates vs Claims
a) Repeatability precision
1 outlier result deleted; N=24

## Slide 68
Comparing Estimates vs Claims
a) Repeatability precision
1 outlier result deleted; N=24

## Slide 69
Comparing Estimates vs Claims
a) Repeatability precision

## Slide 70
Comparing Estimates vs Claims
a) Repeatability precision

## Slide 71
Comparing Estimates vs Claims
a) Repeatability precision

## Slide 72
Comparing Estimates vs Claims
a) Repeatability precision

## Slide 73
Comparing Estimates vs Claims
a) Repeatability precision

## Slide 74
Comparing Estimates vs Claims
a) Repeatability precision

## Slide 75
Comparing Estimates vs Claims
b) Within-lab precision

## Slide 76
Comparing Estimates vs Claims
b) Within-lab precision

## Slide 77
Comparing Estimates vs Claims
b) Within-lab precision

## Slide 78
Comparing Estimates vs Claims
b) Within-lab precision

## Slide 79
Comparing Estimates vs Claims
b) Within-lab precision

## Slide 80
Considering deleting one result as statistical outlier:
Repeatability & within-lab precision at all 3 levels were confirmed
This study is accepted as a whole: The estimates are generally consistent with the manufacture’s claims
There is no problem with middle and high levels; but it’s recommended performance at low level be monitored carefully
Interpretation

## Slide 86
Part 2

## Slide 87: Statistical basis
H0: Bias is zero; lab’s Mean is equal to TV
HA: Bias isn’t zero; lab’s mean differs from TV

## Slide 88: Statistical basis
Reference Material (MR) with Target Values
Assaying RM to estimate user Mean
Mean
TV
BIAS = Mean - TV
H0: Bias is zero; lab’s Mean is equal to TV
HA: Bias isn’t zero; lab’s mean differs from TV

## Slide 89: Statistical basis
Reference Material (MR) with Target Values
Assaying RM to estimate user Mean
Mean
TV
Standard uncertainty of RM, uRM
Standard uncertainty of Mean, uMean

## Slide 90: Statistical basis
ubias =
Standard uncertainty of Bias

## Slide 91: Statistical basis
ubias =
Ubias = m * ubias
Standard uncertainty of Bias
Expanded uncertainty of Bias
TV

## Slide 92: Statistical basis
ubias =
Ubias = m * ubias
Standard uncertainty of Bias
Expanded uncertainty of Bias
TV
TV - Ubias
TV + Ubias

## Slide 93: Statistical basis
ubias =
Ubias = m * ubias
Standard uncertainty of Bias
Expanded uncertainty of Bias
TV - Ubias
TV + Ubias
Verification Interval
H0: Bias is zero; lab’s Mean is equal to TV
HA: Bias isn’t zero; lab’s mean differs from TV
Lab Mean within VI → Bias Statistically Not Significant
Lab Mean out of VI → Bias Statistically Significant

## Slide 94: Overview
Evaluating BIAS via assaying reference materials (RM); materials with assigned target values
Allowable bias, Ba, specified before study
Bias may be “statistically” significant but “clinically” acceptable
, SR and SWL from “Precision Verification” study are used

## Slide 95
Study’s purpose
Depending on the reference material (RM):
Confirm consistency with target value of a CRM
Confirm consistency with TV of EQA/PT
Recovery of an expected value (Spiked material)

## Slide 96
Uncertainty of the target value, uRM, must be known
Depending on material, N & DF may be necessary
Standard Reference (CRM) - JCTLM, IFCC, NIST, WHO,…
uRM presented in certificate
dfRM = ∞
May not appropriate for some lab: Accessibility, commutability, cost, stability
Spiked sample - e.g. drugs
Practically uRM is considered zero and dfRM considered ∞
Specifications of RMs

## Slide 97
EQA materials – Preferably unaltered patient samples
TV = Group Mean
uRM = SDGroup/
dfRM = nLab – 1
PT materials
TV, uRM & dfRM as EQA
nLab > 10; preferably > 20
Some labs provide more results
QC & Calibration Verification materials
uRM assumed zero
dfRM assumed infinity
Unrealistic assumptions

## Slide 98
Conventional TVs
uRM = 0
dfRM = ∞
Patient samples tested in one or more central labs
Patient samples sent to a lab traceable to a reference hierarchy
Previously collected QC data
The observed difference reflects “TRUENESS” if a high quality method is used for assigning TV

## Slide 99: User Mean, its DF & uncertainty
Mean = (from precision verification study)
u can be calculated from 3 ways:
Using precision estimates (sR and sWL)
Precision is verified, use σR and σWL in the above formula
An equivalent equation u = sWL /
df = nRun - 1
u

## Slide 100
TV
uTV
u
Bias = - TV
ubias =
Ubias =
t-value for Padj

## Slide 101: Determine multiplier “m” (to expand uncertainty)
Combined DF (Satterthwaite’s approximation)
2
Adjust Probability for nSam
Padj = 1 – [0.05/(2*nSam)]
Determine “m” from t distribution
m = t(Padj, dfc)
nSam = 1 → 0.975
nSam = 2 → 0.9875
nSam = 3 → 0.9917

## Slide 102
TV
uTV
u
Bias = - TV
ubias =

## Slide 103
m
TV
uTV
u
Bias = - TV
ubias =
Ubias = m * ubias

## Slide 104
m
TV
uTV
u
Bias = - TV
ubias =
Ubias = m * ubias
VI = TV ± Ubias
Verification Interval

## Slide 105
m
TV
uTV
u
Bias = - TV
ubias =
Ubias = m * ubias
VI = TV ± Ubias
TV
+ UBias
- UBias
Statistically Significant Bias

## Slide 106: Interpretation
1. UBias < Allowable Bias?
NO: Cannot detect bias as large as Ba
Increase Run and/or Replicates
YES: Enough power to detect Ba
Evaluate “Bias” (2nd question)

## Slide 107
2. “VI” includes “”?
YES: Bias is NOT statistically significant
NO: Statistically significant bias.
Evaluate “Clinical Importance” (3rd question)
3. Bias < Allowable Bias?
YES: Clinically acceptable bias
NO: Clinically unacceptable bias; reject method

## Slide 108
EXAMPLES

## Slide 109: Estimating bias for a Ferritin assay
Purpose: Confirm consistency with TV of EQA at 3 levels
Reference material (Normal level)
TV = Group Mean = 142.5 μg/L
nLab = 43
SDGroup= 4.5 μg/L
uRM = SDGroup/ = 0.69 μg/L
dfRM = nLab – 1 = 42

## Slide 110
Laboratory
Allowable Bias = 10% → 142.5 * 10% = 14.25
nRun = 5
nRep = 5
= 140.1
SR = 1.78
SWL = 2.40
u = 0.8
df = nRun – 1 = 4
u

## Slide 111
Lab
= 140.1
u = 0.8
df = 4
RM
TV = 142.5
uRM = 0.69
dfRM = 42
ubias =
=
ubias = 1.06
2
dfC = 9
Padj = 1 – [0.05/(2*nSam)]
= 1 – [0.05/(2*3)]
Padj = 0.9917
m = t(Padj, dfc)
= t(0.9917, 9)
m = 2.94
Ubias = m * ubias
= 2.94 * 1.06
Ubias = 3.12

## Slide 112
Verification Interval
m
TV
uTV
u
Bias = - TV
ubias =
Ubias = m * ubias
VI = TV ± Ubias
= 142.5 ± 3.125
VI = [139.4, 145.5]
142.5
Ba = 14.2
TV = 142.5
Ubias = 3.125
= 140.1
140.1
139.4
145.5
Ubias < Ba?
3.125 < 14.25
Enough data to detect bias = 14.25
H0, i.e. absence of bias, CANNOT be rejected
Bias in NOT statistically significant

## Slide 126: Successful completion of EP15
Laboratory demonstrates:
Consistency of precision in lab with manufacture’s claim
No significant or acceptable bias relative to the selected RM

## Slide 127
THANKS!